Annex VII

### BID/PROPOSAL SUBMISSION FORM

To: The procuring entity

Dear Sir / Madam,

Having examined the Bidding Documents, the receipt of which is hereby duly acknowledged, we, the undersigned, offer to supply and deliver [description of goods] in conformity with the said bidding documents for the sum of [total bid amount in words and figures] as may be ascertained in accordance with the Price Schedule attached herewith and made part of this Bid.

We undertake, if our Bid is accepted, to deliver the goods in accordance with the delivery schedule specified in the Schedule of Requirements.

We agree to abide by this Bid for a period of [number] days from the date fixed for opening of Bids in the Invitation to Bid, and it shall remain binding upon us and may be accepted at any time before the expiration of that period.

We understand that you are not bound to accept any Bid you may receive.

Dated this . . . . .day of . . . . .[year].

. . . . . . . . . . . . . . . . . . . . . . . . .			. . . . . . . . . . . . . . . . . . . . . .

Signature 						[in the capacity of]

Duly authorised to sign the Bid for and on behalf of . . . . . . . . . . . . . . . . . . . . . . . . . . . .

Annex VIII

PRICE SCHEDULE

- The Price Schedule must provide a detailed cost breakdown for each item.

- Technical descriptions for each proposed item must provide sufficient detail to allow the Purchaser to determine compliance of Bid with specifications as per Schedule of Requirements and Technical Specifications of this ITB.

- Estimated weight/volume of the consignment must be part of the documentation submitted.

- All prices/rates quoted must be exclusive of all taxes, since the United Nations, including its subsidiary organs, is exempt from taxes.

- The format shown on the following pages should be used in preparing the Price Schedule. The format uses a specific structure which may or may not be applicable but are indicated to serve as examples.

- In addition to the hard copy, if possible please provide also the information in electronic format.

| Name of Bidder:………………………….. |
| --- |
| LOT 1 - Pathology lab equipment |
| Item nr | Item | Description | Unit | Unit Price | Quantity Required | Total Price per item |
| 1 | Multistainer with coverslipper |  |  |  | 1 |  |
| 1.1 | Slides |  |  |  | 15,000 |  |
| 1.2 | Coverslip |  |  |  | 15,000 |  |
| 2 | Semiautomatic microtome |  |  |  | 1 |  |
| 2.1 | Blades |  |  |  | 5 |  |
| 3 | Histobath |  |  |  | 2 |  |
| 4 | Histoplate |  |  |  | 2 |  |
| 5 | Incubator for histology |  |  |  | 1 |  |
| 6 | Trinocular Microscope |  |  |  | 3 |  |
| 7 | Polarized light Microscope |  |  |  | 1 |  |
| 8 | Phase contrast Microscope |  |  |  | 1 |  |
| 9 | Trinocular Microscope with immersion |  |  |  | 6 |  |
| 10 | Training |  | days |  | 5 |  |
| 11 | Warranty and maintenance services |  | years |  | 2 |  |
| 12 | Metrological certifications |  |  |  |  |  |
| 13 | Delivery and installation |  |  |  |  |  |
|  | GRAND TOTAL |  |  |  |  |  |
| LOT 2 - Analytical equipment |
| Item nr | Item | Description | Unit | Unit Price | Quantity Required | Total Price per item |
| 1 | Atomic absorption spectrometer |  |  |  | 1 |  |
| 2 | Microwave Digestion System |  |  |  | 1 |  |
| 3 | High Performance Liquid Chromatograph |  |  |  | 1 |  |
| 4 | Gas Chromatograph |  |  |  | 1 |  |
| 4.1 | Vials |  |  |  | 300 |  |
| 4.2 | Vial caps |  |  |  | 300 |  |
| 4.3 | Ferrules |  |  |  | 10 |  |
| 4.4 | Nuts |  |  |  | 4 |  |
| 4.5 | Liner |  |  |  | 1 |  |
| 5 | Drug detection system |  |  |  | 1 |  |
| 6 | Semiautomatic preparing system |  |  |  | 1 |  |
| 7 | Water purification system |  |  |  | 1 |  |
| 8 | N2 generator |  |  |  | 1 |  |
| 9 | Rotary evaporator |  |  |  | 1 |  |
| 10 | Training |  | days |  | 5 |  |
| 11 | Warranty and maintenance |  | years |  | 2 |  |
| 12 | Metrological certifications |  |  |  |  |  |
| 13 | Delivery and installation services |  |  |  |  |  |
|  | GRAND TOTAL |  |  |  |  |  |
| LOT 3 - General lab equipment |
| Item nr | Item | Description | Unit | Unit Price | Quantity Required | Total Price per item |
| 1 | Electric stirrer |  |  |  | 1 |  |
| 2 | Laboratory centrifuge |  |  |  | 2 |  |
| 3 | Ultrasonic bath |  |  |  | 1 |  |
| 4 | Laboratory incubator |  |  |  | 1 |  |
| 5 | Colposcope |  |  |  | 1 |  |
| 6 | UV box sterilizer |  |  |  | 1 |  |
| 7 | Roentgen film developing system |  |  |  | 1 |  |
| 8 | Gynecological couch |  |  |  | 1 |  |
| 9 | Refrigerator for biological samples |  |  |  | 29 |  |
| 10 | Training |  | days |  | 3 |  |
| 11 | Warranty and maintenance services |  | years |  | 2 |  |
| 12 | Metrological certifications |  |  |  |  |  |
| 13 | Delivery and installation |  |  |  |  |  |
|  | GRAND TOTAL |  |  |  |  |  |

- *Unit price should be based on Incoterms used i.e. FOB/FCA/C&P/C&F/DU as the case may be.

- Note: In case of discrepancy between unit price and total, the unit price shall prevail.

Signature of Bidder ………………………………………..

Annex IX

TECHNICAL COMPLIANCE FORM

- The Technical Compliance Form must provide a detailed specification breakdown for each item.

- Technical descriptions for each proposed item must provide sufficient detail to allow the Purchaser to determine compliance of Bid with specifications as per Schedule of Requirements and Technical Specifications of this ITB.

- The format shown on the following pages should be used in preparing the Technical Compliance Form. The format shall indicate individual specifications for each item and/or any other specifications you may find relevant for this item.

| Name of Bidder:………………………….. |
| --- |
| LOT 1 - Pathology lab equipment |
| Item | Item Description | Minimum required specification | Proposed specification |
| 1 | Multistainer with coverslipper | Proposed model: ... |
|  | Stainer Type | Robotic- fully automatic |  |
|  | Number of cuvettes | At least 40 |  |
|  | Cuvette volume | Between 350 and 450 ml |  |
|  | Drying station | 1 |  |
|  | Loading station | 4 |  |
|  | Unloading station | 4 |  |
|  | Water bath | At least 4 |  |
|  | Number of protocols | At least 18 |  |
|  | Protocols | Different basket to be able to run different protocol |  |
|  | Maxim capacity | At least 550 slides / hour |  |
|  | Safety | Fully enclosed, carbon filter |  |
|  | Display | Color LCD |  |
|  | Memory card | Yes, 4 GB capacity |  |
|  | Combination with cover slipper | Inline |  |
|  | Dispensing volume | Up to 220µl |  |
|  | Slide 26 x76mm capacity | 120 pcs |  |
|  | Coverglass capacity | 350pcs |  |
|  | Consumables |  |
|  | Slides | 15,000 pcs |  |
|  | Coverslip | 15,000 pcs |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 2 | Semiautomatic microtome | Proposed model: ... |
|  | Safety lock on hand wheel in every position | Yes |  |
|  | Section thickness range | 0 – 100 µm 0 – 2 in 0,5 µm 2 – 20 in 1 µm 20 – 50 in 2 µm 50 – 100 in 5 µm |  |
|  | Trimming thickness setting | 0 – 300 µm free selectable |  |
|  | Vertical stroke | 60mm |  |
|  | Display of cutting thickness and thickness during trimming | LCD |  |
|  | Specimen orientation | x-axis 8° y-axis 8° z-axis 360° |  |
|  | Specimen retraction | 0 – 200 µm free selectable |  |
|  | Ready for use with all needed accessories | Yes |  |
|  | Microtome slides included | 10 pcs |  |
|  | Consumables |  |
|  | Blades | 5 pcs |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 3 | Histobath | Proposed model: ... |
|  | Heating capacity | Yes |  |
|  | Lid | Stainless steel |  |
|  | Interior | Teflon coated |  |
|  | Over temperature protection | Yes |  |
|  | Electronic temperature regulator | Yes |  |
|  | Temperature display | Yes |  |
|  | Temperature selection | Up to 100°C |  |
|  | Volume | At least 3 liters |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 4 | Histoplate | Proposed model: ... |
|  | Heated plate | Teflon coated |  |
|  | Temperature selectable | Up to 100°C |  |
|  | Electronic temperature controller with digital display | Yes |  |
|  | Dimensions of heated plate | At least 300x 250mm |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 5 | Incubator for histology | Proposed model: ... |
|  | Volume | 110-120 liters |  |
|  | Display | Color LCD |  |
|  | Exterior | Stainless steel |  |
|  | Interior | Stainless steel |  |
|  | Forced air convection | Yes |  |
|  | Fan speed regulation | From 0% to 100% |  |
|  | Temperature | 5°C above ambient-to 100°C |  |
|  | Temperature resolution | 0,1°C |  |
|  | Temperature fluctuation | 0,2°C |  |
|  | Over temperature protection | Class 3.1 |  |
|  | Temperature-time profile | At least 6 segments |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 6 | Trinocular Microscope | Proposed model: ... |
|  | Illumination | White LED or halogen 35W with light intensity control |  |
|  | Observation | Bright field |  |
|  | Revolver | Quadruple |  |
|  | Head | Trinocular inclined 30° and rotatable 360° |  |
|  | Interpupillary adjustment | 55-75mm |  |
|  | Diopter adjustment | On both eyepieces |  |
|  | Condenser | Abbe type. 1,25 with centering device |  |
|  | Objectives | 4X, 10X, 40X, 60X anti-fungus treated |  |
|  | Stage | Size 160x140mm, specimen holder for 2 slides |  |
|  | Dust cover | Yes |  |
|  | Digital Camera | At least 5 Mp, C-mount type with all needed accessories for connection to the microscope and to the PC, software , live view possibility |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 7 | Polarized light Microscope | Proposed model: ... |
|  | Illumination | White LED or halogen 35W with light intensity control |  |
|  | Observation | Bright field, polarized light |  |
|  | Revolver | Quintuple with centering system |  |
|  | Head | Trinocular inclined 30° and rotatable 360° |  |
|  | Interpupillary adjustment | 55-75mm |  |
|  | Diopter adjustment | On both eyepieces |  |
|  | Condenser | Rotatable polarizer with graduated scale |  |
|  | Objectives | 4X, 10X, 25X, 40X and 63X anti-fungus treated |  |
|  | Stage | Size 160x140mm, specimen holder for 2 slides |  |
|  | Dust cover | Yes |  |
|  | Digital Camera | At least 5 Mp, C-mount type with all needed accessories for connection to the microscope and to the PC, software , live view possibility |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 8 | Phase contrast Microscope | Proposed model: ... |
|  | Illumination | White LED or halogen 35W with light intensity control |  |
|  | Observation | Bright field, dark filed, phase contrast |  |
|  | Revolver | Quintuple |  |
|  | Head | Trinocular inclined 30° and rotatable 360° |  |
|  | Interpupillary adjustment | 55-75mm |  |
|  | Diopter adjustment | On both eyepieces |  |
|  | Condenser | Phase contrast condenser, bright field, dark field |  |
|  | Objectives | 4X, 10XPh, 40 XPh, 100XPh anti-fungus treated |  |
|  | Stage | Size 160x140mm, specimen holder for 2 slides |  |
|  | Dust cover | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 9 | Trinocular Microscope with immersion | Proposed model: ... |
|  | Illumination | White LED or halogen 35W with light intensity control |  |
|  | Observation | Bright field |  |
|  | Revolver | Quadruple |  |
|  | Head | Trinocular inclined 30° and rotatable 360° |  |
|  | Interpupillary adjustment | 55-75mm |  |
|  | Diopter adjustment | On both eyepieces |  |
|  | Condenser | Abbe type. 1,25 with centering device |  |
|  | Objectives | 4X, 10X, 40 X, 100X anti-fungus treated |  |
|  | Stage | Size 160x140mm, specimen holder for 2 slides |  |
|  | Dust cover | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| Common requirements for Lot 1 | Compliance with common requirements |
| Data Connection | When required, data connection (Interface to personal computer) may be RS232C, RS423, IEEE-488, LAN or USB. A 5 m appropriate cable with standard PC connector will be supplied. |  |
| Software | Unless otherwise specified, any special software needed to operate the item(s) proposed should be supplied. The software should operate on a standard Windows compatible PC along with the data cable mentioned above. If there is a specific requirement for the PC, it should be mentioned in the offer. |  |
| Environmental conditions | Unless otherwise specified, equipment to be procured will be used in controlled environment with temperature 20±3°C, and relative humidity 55±15 %. Electricity supply is 230±30 V and 50 Hz. |  |
| Quality | Good Laboratory Practice (GLP), CE Certificate, ISO 9001 Certificate |  |

| Name of Bidder:………………………….. |
| --- |
| LOT 2 - Analytical equipment |
| Item | Item Description | Minimum required specification | Proposed specification |
| 1 | Atomic absorption spectrometer | Proposed model: ... |
|  | Wavelength Range | 190nm – 900nm |  |
|  | Monochromator | Czerny-Turner Configuration |  |
|  | Spectral Bandwidth | 0.1nm, 0.2nm, 0.4nm, 1.0nm, 2.0nm, (software selectable). |  |
|  | Wavelength Accuracy | ± 0.15nm |  |
|  | Wavelength Reproducibility | < 0.05nm |  |
|  | Resolution | 0.2nm ± 0.02nm |  |
|  | Baseline Stability | 0.005A/30min |  |
|  | Sensitivity (Cu) | 2 µg/ml Absorption > 0.28A (flame) |  |
|  | Detection Limit | Cu < 0.004 µg/ml (flame) Cd < 0.4 x 10-12 g (graphite furnace) |  |
|  | Repeatability | Cu < 0.7% (Air/Acetylene flame) Ba < 1.0% ( Nitrous oxide/Acetylene flame) Cu < 2.0% Cd < 2.0% (Graphite Furnace) |  |
|  | Position Adjustment | Automatic changeover |  |
|  | Safety Functions | Burner Identification, Flame Sensor, Gas leak Sensor, Low Gas Pressure Sensor, Drain Trap Sensor, Power Loss Protection,Circulation Water (graphite), Over Temperature Sensor (graphite). For Hg analise the system must be supplied with protection capsule for operator safety. |  |
|  | Hydride generator | Electro-heating, continuous flow peristaltic pump with speed control, high efficiency mixing section and gas-liquid separation |  |
|  | Autosampler | For flame and graphit |  |
|  | Monoelement lamps | Cu, Zn, Pb, As, Hg, P, Cr, Mn, Cd, Tl, Ba, Sb |  |
|  | Computer | Processor 3Ghz, 2GB RAM, 500GB HDD, 19” LCD monitor, Windows 7 license |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 2 | Microwave Digestion System |  |
|  | System with high-pressure capabilities to enhance digestion quality while decreasing digestion time | Yes |  |
|  | Operation and full computer control of temperature and power | Yes |  |
|  | Software with direct control and sensing of temperature with real-time graphic display, automatic data storage, unlimited method storage, and complete record of instrument parameters | Yes |  |
|  | External PC with software control | Yes |  |
|  | Top-loading with safety interlocks | Yes |  |
|  | Fume exhaust module | Yes |  |
|  | Teflon lining of oven cavity | Yes |  |
|  | Continuous temperature monitoring and control | Yes |  |
|  | Release of acid fumes into collection reservoir in vessel carousel | Yes |  |
|  | Automatic pressure release mechanism in oven cavity | Yes |  |
|  | Number of Vessels | 6 |  |
|  | Operating Pressure | 400psi |  |
|  | Operating Temperature | 250°C |  |
|  | Temperature Sensor | 1 |  |
|  | Vessel Volume | 60mL |  |
|  | Microwave Power | 1200W |  |
|  | Magnetron Frequency | 2450 Mhz |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 3 | High Performance Liquid Chromatograph |  |
|  | Quaternary Pump |  |
|  | Hydraulic system | Dual piston in series pump with servo-controlled variable stroke drive, floating pistons |  |
|  | Settable flow range | Set points 0.001 - 10 mL/min, in 0.001 mL/min increments |  |
|  | Flow range | Up to 10.0 mL/min |  |
|  | Flow precision | <= 0.05 % RSD or <= 0.01 min SD whichever is greater |  |
|  | Flow accuracy | ± 0,1 % |  |
|  | Pressure operating range | Operating range up to 60 MPa (600 bar, 8700 psi) up to 10 mL/min |  |
|  | Pressure pulsation | Typically < 0.2 MPa (2 bar) or <1% whichever is greater |  |
|  | Compressibility compensation | User selectable, based on mobile phase compressibility |  |
|  | Gradient formation | Low pressure quaternary mixing/gradient capability |  |
|  | Composition precision | < 0.15 % SD at 1 mL/min |  |
|  | Integrated degassing unit | Number of channels: 4 Internal volume per channel: 1.5 mL |  |
|  | Control | By software |  |
|  | Local control | Touch screen |  |
|  | Communications | Controller-area network (CAN), RS-232C, APG Remote: ready, start, stop and shut-down signals, LAN optional |  |
|  | Safety and maintenance | Extensive diagnostics, error detection and display, leak detection, safe leak handling, leak output signal for shutdown of the pumping system. Low voltage in major maintenance areas. |  |
|  | Degasser | 4 channel, integrated |  |
|  | Injector |  |
|  | Injection methods | Full-loop and partial-loop injections |  |
|  | Injection loop | 20 μL sample loop |  |
|  | Column Thermostat |  |
|  | Temperature range | 6°C… 80 °C |  |
|  | Temperature stability | ± 0.1 °C |  |
|  | Temperature accuracy | ± 0.5 °C |  |
|  | Column capacity | Four 30 cm columns |  |
|  | Heat-up/cool-down time | 15 min from 20 °C to 50 °C 15 min from 50 °C to 20 °C |  |
|  | DAD Detector |  |
|  | Detector type | 1024-element diode array, single-beam |  |
|  | Light source | Deuterium lamp, tungsten lamp, temperature control for both lamps |  |
|  | Number of signals | 8 |  |
|  | Maximum sampling rate | 100 Hz (both spectra and signals) |  |
|  | Short-term noise | Wide slit: < ±8 μAU at 254 nm Narrow slit:< ±10 μAU at 254 nm response time 2 s (according to ASTM time constant ~1 s), 4 nm bandwidth, water at 1.0 mL/min |  |
|  | Drift | < 1 mAU/h (typically < 0.5 mAU/h) at 254 and 520 nm with water at 1.0 mL/min |  |
|  | Linearity | < 3% RSD and corr. coeff. > 0.9995 up to 1.5 AU, typically < 5% RSD and corr. coeff. > 0.999 up to 2.0 AU |  |
|  | Wavelength range | 190-800 nm |  |
|  | Wavelength accuracy | ±1.0 nm |  |
|  | Slit width | Narrow or wide |  |
|  | Pixel Resolution | <1 nm |  |
|  | Flow cells | Standard (Stainless Steel) 13-µL volume, 10 mm-cell path length, 120 bar (1740 psi) pressure maximum |  |
|  | RI Detector |  |
|  | Detection type | Deflection method |  |
|  | Short-term noise | <2.5 × 10-9 RIU |  |
|  | Drift | < 200 x 10-9 RIU/h |  |
|  | Refractive index range | 1.00–1.75, calibrated |  |
|  | Range | 0.25 to 512 µRIU |  |
|  | Linearity | 600 µRIU |  |
|  | Flow cell | 8 µL, 5 bar pressure maximum |  |
|  | Temperature control | 30 °C to 50 °C |  |
|  | Maximum flow rate | 10 ml/min |  |
|  | Time programmable | Polarity, peak width |  |
|  | Response | 0.1, 0.25, 0.5, 1.0, 1.5, 2.0, 3.0, 6.0 s |  |
|  | Zero adjustment | Optical and electronic autozero |  |
|  | Analog output | Recorder/integrator: 10 mV or 1 V |  |
|  | Communications | RS-232, remote ready, start and shutdown signals |  |
|  | Software | Yes |  |
|  | Accessories |  |
|  | Syringe | 3 pcs of 100 µL |  |
|  | Column | C18, 250 mm, ID 4.6, 5 µm |  |
|  | Computer | Processor 3Ghz, 2GB RAM,500GB HDD, 19” LCD monitor, Windows 7 license |  |
|  | Laser printer | 10 pages/min, B/W |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 4 | Gas Chromatograph |  |
|  | GC/FID/TOF-MS/HS |  |
|  | Number of injectors to be installed simultaneously in a chromatograph | 3 |  |
|  | Number of detectors to be installed simultaneously in a chromatograph | 3 |  |
|  | Column oven |  |
|  | Volume | 12-13 L |  |
|  | Operating temperature | 5 above ambient to 500 0C |  |
|  | Temperature resolution | 0.1°C |  |
|  | Number of programmable steps of oven temperature | 25 |  |
|  | Max Heating Ramp Rate | 140°C/min, 0.1 °C/min resolution |  |
|  | Isothermal time | 0.00 to 999.00 min, 0.01 min resolution |  |
|  | Injector (2 units) | Split/Splitless |  |
|  | Temperature | Up to 450°C, 1°C resolution |  |
|  | Injector Carrier gas max total flow | 1000 mL/min |  |
|  | Injector Pressure range | 0 - 120 psi |  |
|  | Control type by flow lines of carrier-gases | Digital flow controller |  |
|  | Detector | FID |  |
|  | Temperature | Up to 450°C, 1°C resolution |  |
|  | Linear dynamic range | > 107 |  |
|  | Minimum Detectable Level | 1.8 pg Carbon/sec |  |
|  | Data Acquisition Rate | 300 Hz |  |
|  | Flame Ignition | Automatic, software controlled |  |
|  | Detector | TOF-MS |  |
|  | Mass Range | 2-1024 amu |  |
|  | Scan Rate | 30,000 spectra/s (up to 1,000 Hz written to disk) equivalent to 1,024,000 u/sec |  |
|  | Acquisition speed | Up to 1000 spectra/sec full scan |  |
|  | Resolution | 2000 on m/z/219 |  |
|  | Sensitivity | 1pg/uL OFN S/N > 150:1 (RMS) in standard operating conditions; S/N>500:1 (RMS in high sensitivity mode full scan |  |
|  | Linearity | 105 |  |
|  | Tune Modes | AutoTune and Manual Tune |  |
|  | Ion Source Mode | EI - Electronic Impact |  |
|  | Ion Source Type | Inert EI source |  |
|  | Source Temperature | Up to 220°C |  |
|  | Heating rate | Up to 140° C / min |  |
|  | Detector | Electron Multiplier |  |
|  | Filaments | Dual filament |  |
|  | Transfer Line Temperature | Up to 300°C |  |
|  | Vacuum System | Integrated oil free baking membrane pump @ 0.2m3/h Two Turbo-molecular Pumps (250 L/s and 70 L/s) |  |
|  | MS Library | NIST mass spectral library - the most updated version available on market |  |
|  | Head Space |  |
|  | Tray capacity | 40-50 vials |  |
|  | Samples | Glass vials 20 ml |  |
|  | Sample path | Silco-treated stainless steel |  |
|  | Vial movement | X -Y robotic arm and gripper |  |
|  | Oven |  |
|  | Capacity | 6 vials |  |
|  | Heating temperature | 15°C above ambient up to 200°C, 1°C resolution |  |
|  | Heating time | From 1 to 300 min, 0.1 min resolution |  |
|  | Heating | Electrical with forced air |  |
|  | Shaking | No, soft or fast regime |  |
|  | Manifold |  |
|  | Temperature | 15°C above ambient up to 200°C, 1°C resolution |  |
|  | Heating | Electrical |  |
|  | Sampling valve | 6 ports, electrically actuated |  |
|  | Sampling loop | 1 ml standard, 3 ml available |  |
|  | Press and vent valves | On - off solenoid valve |  |
|  | Transfer line |  |
|  | Temperature | 15°C above ambient up to 220°C, 1°C resolution |  |
|  | Heating | Electrical |  |
|  | Needle | Removable, standard and for on-column injector |  |
|  | Dimensions | 1 m length, 0.8 mm i.d |  |
|  | Column | 30m, diam 0.53mm, particle 3.00u |  |
|  | Column | 25mm, diam 0.25mm, particle 0.25um |  |
|  | Computer | Processor 3Ghz, 2GB RAM,500GB HDD, 19” LCD monitor, Windows 7 license |  |
|  | Consumables |  |
|  | Vials | 300 pcs |  |
|  | Vial caps | 300 pcs |  |
|  | Ferrules | 10 pcs |  |
|  | Nuts | 4 pcs |  |
|  | Liner | 1 pcs |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 5 | Drug detection system- similar to “Toxilab AB Plus” |  |
|  | Method | Thin layer chromatography |  |
|  | Screening method for polydrug specimens | Yes |  |
|  | Supplied with database for drugs photo-documentation | Yes |  |
|  | Supplied with | 100 tubes, 100 disks blank, 50 disks B1 type, 50 disks B2 type, 50 disks B3 type, 50 disks B4 type, 6 sets of reagents concentrated B1, B2 and B3 or similar |  |
|  | Ready to use system | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 6 | Semiautomatic preparing system |  |
|  | Application | Sample clean up Sample preparation for pesticide analysis, active pharmaceutical compounds, drugs. |  |
|  | Type | Gel permeation chromatography |  |
|  | Injection | Manual |  |
|  | Sample processing | Automatically |  |
|  | Number of columns | One, upgradable to four |  |
|  | Display | Touch screen |  |
|  | PC connection | Yes |  |
|  | Column refill on site | Yes |  |
|  | User protection | The column to be encased in a aluminum sheath |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 7 | Water purification system |  |
|  | Application | HPLC, GC/MS, AA |  |
|  | Water type | Type 1 according to ISO3696 |  |
|  | Resistivity | 18.2MΩ-cm |  |
|  | UV lamp | Yes |  |
|  | TOC, ppb | 1-5 |  |
|  | Bacteria, CFU/ml | <1 |  |
|  | Output | 15 L/ day |  |
|  | Internal non-contaminating reservoir | At least 6 liters |  |
|  | Feed water | Distilled water |  |
|  | Standby function | Yes |  |
|  | Final filter | 0.2 micron |  |
|  | Set of filters | Yes. Complete package depending on configuration |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 8 | N2 generator |  |
|  | Application | Gas chromatograph connection |  |
|  | Purity | 99.9995% |  |
|  | N2 Flow rate | 1 L/min |  |
|  | Air Flow rate | 1,5 L/min |  |
|  | Outlet pressure | 5,5 bars |  |
|  | Integrated air compressor | Yes |  |
|  | PSA technology | Yes |  |
|  | Touch screen | Yes |  |
|  | Flow produced indication | Yes |  |
|  | Energy saving mode | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 9 | Rotary evaporator |  |
|  | Application | Sample preparation |  |
|  | Rotation speed | 20-270rpm |  |
|  | Speed set | Digital set of rotation speed with display |  |
|  | Lift | Motor lift 18mm/sec |  |
|  | Heating bath temperature range | 20-180°C |  |
|  | Over heat protection | Yes |  |
|  | Heating bath volume | 4-5 liters |  |
|  | Digital display of vapor temperature | Yes |  |
|  | Timer | Automatic cut off at the end of distillation |  |
|  | Remote control | Yes |  |
|  | Protection class | IP 20 |  |
|  | Supplied with G3 set | Yes |  |
|  | Modular design with waterbath, vacuum pump and operating panel | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| Common requirements for Lot 2 | Compliance with common requirements |
| Data Connection | When required, data connection (Interface to personal computer) may be RS232C, RS423, IEEE-488, LAN or USB. A 5 m appropriate cable with standard PC connector will be supplied. |  |
| Software | Unless otherwise specified, any special software needed to operate the item(s) proposed should be supplied. The software should operate on a standard Windows compatible PC along with the data cable mentioned above. If there is a specific requirement for the PC, it should be mentioned in the offer. |  |
| Environmental conditions | Unless otherwise specified, equipment to be procured will be used in controlled environment with temperature 20±3°C, and relative humidity 55±15 %. Electricity supply is 230±30 V and 50 Hz |  |
| Consumbales | For Item 1: Atomic absorption spectrometer please quote separately prices for each lamp requested |  |
| Quality | Good Laboratory Practice (GLP), CE Certificate, ISO 9001 Certificate |  |

| Name of Bidder:………………………….. |
| --- |
| LOT 3 - General lab equipment |
| Item | Item Description | Minimum required specification | Proposed specification |
| 1 | Electric stirrer | Proposed model: ... |
|  | Vibration orbit | 5mm |  |
|  | Speed | 2500rpm |  |
|  | Continuous operation or automatic | Yes |  |
|  | Adapter for 10 ml tubes | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 2 | Laboratory centrifuge |  |
|  | Centrifugation speed | 100 rpm to 18000 rpm variable speed |  |
|  | Maximum capacity | At least 1000ml |  |
|  | Rotor | Swing out type, speed 5000 RMP for 4 x 250ml tubes, 32 x 10-15ml tubes with all needed buckets, adapters and tubes |  |
|  | Display | LCD |  |
|  | Memory | At least 100 programs |  |
|  | Timer | 0-99min 59s and continuous |  |
|  | Automatic rotor identification | Yes |  |
|  | Unbalance warning | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 3 | Ultrasonic bath |  |
|  | Capacity | 2,5-3,5 litres |  |
|  | House and cleaning tank | Stainless steel |  |
|  | Temperature range regulated by microprocessor | 30-80°C |  |
|  | Auto degas function | Yes |  |
|  | Display | LED |  |
|  | Working frequency | 25-40 Hz |  |
|  | Cover | Stainless steel |  |
|  | Basket | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 4 | Laboratory incubator |  |
|  | Volume | 55-65 liters |  |
|  | Exterior | Stainless steel |  |
|  | Interior | Stainless steel |  |
|  | Forced air convection | Yes |  |
|  | Fan speed regulation | From 0% to 100% |  |
|  | Temperature | 5°C above ambient - to 100°C |  |
|  | Temperature resolution | 0,1°C |  |
|  | Temperature fluctuation | 0,2°C |  |
|  | Over temperature protection | Class 3.1 |  |
|  | Aver temperature alarm | Yes |  |
|  | Temperature-time profile | At least 6 segments |  |
|  | Display | Color LCD |  |
|  | Network connection | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 5 | Colposcope |  |
|  | Type | Videocolposcope |  |
|  | Autofocus | Up to 30 cm |  |
|  | Electronic green filter | Yes |  |
|  | Continuous zoom | 1-40X |  |
|  | Image processor | Color CCD |  |
|  | Pixels | At least 2 Mpx |  |
|  | Resolution | 825 TV lines |  |
|  | E-flip | Yes |  |
|  | S/N ratio | > 48DB |  |
|  | Light | LED |  |
|  | Color temperature | 7000K |  |
|  | Video out | Yes |  |
|  | Memory card | Yes, 8GB |  |
|  | Photo possibility | Yes- at least 2 Mpx |  |
|  | Safety standards | IEC 60601 |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 6 | UV box sterilizer |  |
|  | Application | Sterilizing surgical instrumentation |  |
|  | UV light | Yes |  |
|  | Box dimensions (W x D x H) | 40 x 40 x 30 cm |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 7 | Roentgen film developing system |  |
|  | Film transport | Continuous roller transport system |  |
|  | Film formats | Up to 35cm width |  |
|  | Processing capacity | 125 films 24 x30cm |  |
|  | Tank capacities | Developer, fixer and washing 5 liters each |  |
|  | Developer temperature | 28-37°C adjustable |  |
|  | Water consumption | No more than 2liter/min when processing |  |
|  | Drain capacity | 7 litres / minute |  |
|  | System protection | IP20 |  |
|  | Noise level | Less than 60dB |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 8 | Gynecological couch |  |
|  | Weight capacity | Up to 180 Kg |  |
|  | 3 segments construction | Yes |  |
|  | Segment independent adjustment | Yes, with manual or automatic adjustment |  |
|  | Surface | Washable |  |
|  | Foot support | Yes |  |
|  | Paper roll holder | Yes |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| 9 | Refrigerator for biological samples |  |
|  | Volume | 100 +/- 15 Litres |  |
|  | Energy class | A+ |  |
|  | Interior | Antibacterial surface easy to clean |  |
|  | Shelves | 2 |  |
|  | Temperature | 2- 6°C |  |
|  |  |  | Please indicate any other specifications you may find relevant for this item … |
| Common requirements for Lot 3 | Compliance with common requirements |
| Data Connection | When required, data connection (Interface to personal computer) may be RS232C, RS423, IEEE-488, LAN or USB. A 5 m appropriate cable with standard PC connector will be supplied. |  |
| Software | Unless otherwise specified, any special software needed to operate the item(s) proposed should be supplied. The software should operate on a standard Windows compatible PC along with the data cable mentioned above. If there is a specific requirement for the PC, it should be mentioned in the offer. |  |
| Environmental conditions | Unless otherwise specified, equipment to be procured will be used in controlled environment with temperature 20±3°C, and relative humidity 55±15 %. Electricity supply is 230±30 V |  |
| Quality | Good Laboratory Practice (GLP), CE Certificate, ISO 9001 Certificate |  |

Annex X

BID/PROPOSAL SECURITY FORM

To:	The procuring entity,

Whereas [name of bidder] (hereinafter called the “Bidder”) has submitted its bid dated [date submission of bid] for the provision of [description of service] (hereinafter called Bid).

KNOW ALL PEOPLE by these presents that WE [name of bank], having our registered office at [address of bank] (hereinafter called the “Bank”), are bound unto [name of Purchaser] (hereinafter called the “Purchaser”) in the sum of for which payment well and truly to be made to the said Purchaser, the Bank binds itself, its successors, and assigns by these presents. Sealed with the Common seal of the said Bank this ………..day of..…….2000.

THE CONDITIONS of this obligation are:

- If the Bidder withdraws its Bid during the period of bid validity specified by the Bidder on the Bid Submission Form; or

- If the Bidder, having been notified of the acceptance of its Bid by the Purchaser during the period of validity of the bid:

- fails or refuses to execute the Contract Form, or

- fails or refuses to furnish the Performance Security, in accordance with the Instructions to Bidders;

we undertake to pay to the Purchaser up to the above amount upon receipt of its first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser will note that the amount claimed by it is due to it, owing to the occurrence of one or both of the two conditions, specifying the occurred condition or conditions.

This guarantee will remain in force up to and including thirty (30) days after the period of validity of the bid, and any demand in respect thereof should reach the Bank not later than the above date.

.............................

Signature of the Bank